# MRI Image Acquisition, Step by Step

How a magnet, three gradient coils and a radio antenna turn water molecules into the numbers this pipeline consumes

## How to read this document

This is the companion to brain.md, which covers what the brain is,

and pipeline_science.md, which covers what the pipeline

does with the images. This document covers the step in between: where the

images come from, and why they have the particular flaws that preprocessing

spends most of its time repairing.

Sections are tagged so you can pick a depth:

(Beginner)* — the idea, with an everyday analogy, no mathematics required.

(Working)* — what a research engineer needs in order to read a protocol

sheet, judge a sidecar, or debug a failed scan.

(Advanced)* — the equations, the second-order effects, and the arguments

people are still having in the literature.

A note on the pictures. Every figure here is computed, not photographed.

There is no participant data in this document and no scanner output either.

The relaxation curves come from the Bloch solutions evaluated at published 3 T

tissue constants; the images come from an analytic brain phantom pushed through

the same signal equations a scanner obeys; k-space, Gibbs ringing and the

distortion pair come from genuine Fourier transforms of that phantom. This is a

deliberate choice: a simulated artefact generated by its real mechanism teaches

the mechanism, whereas a screenshot only teaches the appearance. The script that

produces them is scripts/make_acquisition_figures.py,

and it needs no input data at all.

## Table of contents

Part I — Where the signal comes from

Spin, $B_0$, and why there is any signal at all

Excitation: tipping the magnetisation over

Relaxation: $T_1$, $T_2$, and $T_2^*$

Part II — Turning signal into a picture

Contrast: what TR and TE actually control

Spatial encoding: giving every voxel an address

k-space: where the scanner really writes its data

Resolution, field of view, and aliasing

Part III — Sequences

Spin echo and gradient echo

Echo-planar imaging: the whole plane in one shot

Accelerating: parallel imaging and multiband

The structural scan: MPRAGE and FLAIR

Part IV — Diffusion

Brownian motion and what diffusion MRI measures

The Stejskal–Tanner experiment and the b-value

Directions, shells, and the gradient table

Designing a diffusion protocol

Part V — Everything that goes wrong

Susceptibility distortion and the reversed-PE trick

Eddy currents, motion, and the rest of the artefact zoo

Noise, SNR, and the trade-off triangle

Part VI — From scanner to pipeline

DICOM to BIDS: what must survive the conversion

How acquisition choices reach the connectome

Acquisition QC: what to check before you process

Glossary · References · Reproducing the figures

# Part I — Where the signal comes from

## 1. Spin, B0, and why there is any signal at all

### (Beginner) The compass needle analogy

Every hydrogen nucleus — a single proton, and your body is full of them because

it is full of water — behaves like a tiny compass needle. Left alone, the needles

point in random directions and cancel out perfectly. There is nothing to measure.

Put them in a strong magnetic field and something subtle happens. The needles do

not snap into alignment like compasses on a desk. Thermal jostling is far too

violent for that. Instead, a very slight majority ends up pointing with the

field rather than against it — at 3 T and body temperature, an excess of roughly

one proton in a hundred thousand.

That tiny imbalance is the entire MRI signal. Everything else in this document is

engineering built on top of a one-in-100,000 surplus. It works only because a

cubic millimetre of tissue contains around $10^{19}$ hydrogen nuclei, so even a

minuscule fraction of them is an enormous number.

Figure 1 — (a) Without a field, spins are randomly oriented and there is no net magnetisation; in $B_0$, a slight excess aligns with the field and sums to $M_0$. (b) Individual spins do not sit still — they precess around the field direction like a leaning spinning top. (c) The precession frequency is exactly proportional to field strength, which is why 3 T scanners operate near 128 MHz, squarely in the FM radio band.

### (Working) Larmor precession

A spin in a magnetic field does not simply align; it precesses about the field

direction, sweeping out a cone. The analogy is a spinning top: gravity pulls it

down, but because it is spinning it wobbles around the vertical instead of

falling. The frequency of that wobble is the Larmor frequency:

$$\omega_0 = \gamma B_0 \qquad\text{or}\qquad f_0 = \bar{\gamma} B_0$$

with $\bar{\gamma} = 42.577$ MHz/T for hydrogen. So:

| Field | Larmor frequency |
| --- | --- |
| 1.5 T | 63.9 MHz |
| 3 T | 127.7 MHz |
| 7 T | 298.0 MHz |

Two consequences matter downstream. First, this is a radio frequency, so MRI

hardware is fundamentally radio engineering — hence the shielded room. Second,

because $f_0$ depends on $B_0$ exactly, any imperfection in the field becomes a

frequency error, and a frequency error becomes a position error once we start

encoding position as frequency. That single fact is the root of the susceptibility

distortion in §16.

### (Advanced) Where the one-in-100,000 comes from

The population difference between the parallel and antiparallel states follows

Boltzmann statistics:

$$\frac{\Delta N}{N} \approx \frac{\gamma \hbar B_0}{2 k_B T}$$

At 3 T and 310 K this is about $1.0 \times 10^{-5}$. The net magnetisation is

$$M_0 = \frac{N \gamma^2 \hbar^2 B_0}{4 k_B T}$$

Note $M_0 \propto B_0$: doubling field strength doubles the available

magnetisation, which is the main reason 3 T displaced 1.5 T for research imaging.

The gain is not free — $T_1$ lengthens with field, susceptibility effects worsen

quadratically, and RF power deposition (SAR) rises roughly with $B_0^2$.

## 2. Excitation: tipping the magnetisation over

### (Beginner) Pushing a swing

The net magnetisation $M_0$ points along the main field, and a detector coil sits

in the plane at right angles to it. In that geometry $M_0$ induces no signal

whatsoever: it is static and pointing the wrong way. To measure it, we must knock

it sideways.

The trick is resonance, and the analogy is a playground swing. Shove a swing at

random moments and nothing much happens. Shove it in time with its natural rhythm

and small pushes accumulate into a large arc. Broadcast a radio pulse at

precisely the Larmor frequency and the same thing happens to the magnetisation:

it spirals away from the field axis into the transverse plane, where the coil can

finally hear it.

### (Working) Flip angle and slice selection

Figure 2 — (a) In the rotating frame, a resonant RF pulse spirals the magnetisation down into the transverse plane. (b) Transverse signal follows $\sin\alpha$, while the longitudinal magnetisation left for the next repetition follows $\cos\alpha$ — the two cannot both be maximised. (c) Slice thickness is set by the ratio of RF bandwidth to gradient strength, so a stronger gradient with the same pulse selects a thinner slice.

The flip angle $\alpha$ is how far the pulse tips the magnetisation, set by

the pulse's amplitude and duration. Transverse signal goes as $\sin\alpha$, so 90°

extracts the most signal from a single shot. But it also leaves nothing behind

($\cos 90° = 0$), so you must wait a full $T_1$ before repeating. Fast sequences

therefore use small flip angles: less signal per shot, but the ability to repeat

every few milliseconds more than compensates.

Slice selection is the elegant part. Apply a gradient along $z$ during the RF

pulse, so Larmor frequency varies with position. A pulse containing only a narrow

band of frequencies then excites only the slab whose spins resonate in that band:

$$\Delta z = \frac{\Delta f_{\text{RF}}}{\bar{\gamma} G_z}$$

Move the pulse's centre frequency and you move the slice. Nothing physically

moves; the scanner just listens and speaks at a different pitch.

### (Advanced) Why the pulse shape is a sinc

Selecting a rectangular slab in space means selecting a rectangular band in

frequency, and the Fourier transform of a rectangle is a sinc. That is why RF

pulses look like the ringing waveform in Figure 2. A truly rectangular profile

would need an infinitely long sinc, so real pulses are truncated and apodised,

giving slightly rounded slice edges. Those imperfect edges are one contributor to

cross-talk between adjacent slices, which is why 2-D protocols often acquire

slices in an interleaved order rather than consecutively.

The small-flip-angle regime has a further convenience: for $\alpha \lesssim 30°$

the Bloch equations are approximately linear, so the slice profile really is the

Fourier transform of the pulse envelope. At 90° and 180° that approximation

fails and pulse design becomes a genuinely non-linear problem.

## 3. Relaxation: T1, T2, and T2

### (Beginner) Two independent clocks

Once tipped, the magnetisation returns to equilibrium by two separate processes

that run at the same time but are otherwise unrelated.

$T_1$ — the recovery clock. Magnetisation climbs back up along the field.

The analogy is a struck bell handing its energy to the surroundings: the system

is giving energy back to the tissue lattice.

$T_2$ — the coherence clock. Immediately after the pulse, all the spins

precess in step. They drift out of step because each one feels a slightly

different local field from its neighbours. The analogy is a choir that starts

perfectly together and gradually spreads apart; the sound fades not because

anyone stopped singing but because they stopped agreeing.

Crucially, $T_2$ is always faster than $T_1$. Coherence is lost long before energy

is returned.

### (Working) The numbers that matter at 3 T

Figure 3 — (a) $T_1$ recovery: CSF takes several seconds, white matter under a second. (b) $T_2$ decay, with the faster $T_2^$ shown dashed. (c) Why the spin echo matters: a 180° pulse at TE/2 reverses the static part of the dephasing, so the signal rebuilds into an echo at TE governed by true $T_2$ rather than $T_2^$.

| Tissue | $T_1$ (ms) | $T_2$ (ms) | $T_2^*$ (ms) |
| --- | --- | --- | --- |
| White matter | ~830 | ~80 | ~53 |
| Grey matter | ~1330 | ~110 | ~66 |
| CSF | ~4000 | ~2000 | ~500 |

Two practical readings. CSF's enormous $T_1$ is why it looks black on a

$T_1$-weighted scan — it has barely begun recovering when you sample it — and why

FLAIR needs an inversion time of ~2.5 s to null it. White matter's short $T_1$ is

why it is the brightest tissue on MPRAGE, which is the contrast FreeSurfer relies

on to place the grey/white surface.

**$T_2^*$** is the one that governs diffusion imaging. Real dephasing has two

sources: random molecular interactions (irreversible, giving $T_2$) and static

field inhomogeneity, from imperfect shim and from tissue susceptibility

differences (reversible). Together:

$$\frac{1}{T_2^*} = \frac{1}{T_2} + \gamma \Delta B_{0,\text{inhom}}$$

$T_2^*$ is always shorter, and it is much shorter near air/bone interfaces.

### (Advanced) The spin echo as time reversal

The static part of the dephasing is deterministic: a spin sitting in a field

that is 5 Hz high accumulates phase at a steady 5 Hz. Apply a 180° pulse at time

TE/2 and every accumulated phase flips sign. That same spin now unwinds at 5 Hz

and arrives back at zero phase exactly at time TE. Every spin does, regardless of

its offset, so the signal reappears as an echo.

Hahn's original insight is often explained with the runners analogy: sprinters of

different speeds spread out along a track, but if you tell them all to turn round

at the same moment, they all cross the start line together. Random molecular

tumbling cannot be reversed this way, which is precisely why the echo is governed

by $T_2$ and not $T_2^*$.

# Part II — Turning signal into a picture

## 4. Contrast: what TR and TE actually control

### (Beginner) Two dials

A pulse sequence repeats: excite, wait, read, repeat. Two timings dominate what

the resulting image looks like.

TR (repetition time) — how long between excitations. Short TR does not give

slow-recovering tissue time to catch up, so tissues that differ in $T_1$ end up

at very different brightnesses. TR exposes $T_1$ differences.

TE (echo time) — how long after excitation you listen. Wait longer and

fast-decaying tissue has faded while slow-decaying tissue has not.

TE exposes $T_2$ differences.

Figure 4 — Four familiar contrasts, all synthesised from the same phantom by changing only the timing. The curves below explain the images: at short TR the $T_1$ recovery curves are far apart, and at long TE the $T_2$ decay curves are far apart. Note the CSF curve in red — nearly flat in $T_2$, which is why CSF is brilliant on $T_2$-weighted images and why FLAIR must null it deliberately.

### (Working) The four workhorse contrasts

| Contrast | Timing | Appearance | Used for |
| --- | --- | --- | --- |
| $T_1$-weighted | short TR, short TE | WM bright, GM grey, CSF black | anatomy, segmentation, FreeSurfer |
| $T_2$-weighted | long TR, long TE | CSF bright, oedema bright | pathology, fluid |
| Proton density | long TR, short TE | reflects water content | contrast-neutral reference |
| FLAIR | inversion nulls CSF | CSF black, lesions bright | white matter lesions, WMH |

One honest caveat visible in the figure: a plain spin echo at 3 T gives only

weak grey/white separation, because $T_1$ contrast is intrinsically modest at high

field. That is exactly why structural $T_1$-weighted scans are almost always

inversion-prepared (MPRAGE), which is what Figure 4 shows. If you ever see a

structural scan with disappointing grey/white contrast, check whether inversion

preparation was actually used.

### (Advanced) Signal equations

Spin echo:

$$S = \rho \left(1 - e^{-TR/T_1}\right) e^{-TE/T_2}$$

Spoiled gradient echo, with flip angle $\alpha$:

$$S = \rho \sin\alpha \, \frac{1 - e^{-TR/T_1}}{1 - e^{-TR/T_1}\cos\alpha} \, e^{-TE/T_2^*}$$

Differentiating the second with respect to $\alpha$ gives the Ernst angle,

$\alpha_E = \arccos(e^{-TR/T_1})$, the flip angle maximising signal at a given TR.

Short-TR sequences run near it.

Inversion recovery:

$$S = \rho \left|1 - 2e^{-TI/T_1} + e^{-TR/T_1}\right| e^{-TE/T_2}$$

Setting the bracket to zero nulls a tissue: $TI_{\text{null}} \approx T_1 \ln 2$

for long TR. For CSF at 3 T that is $4000 \times 0.693 \approx 2770$ ms, which is

where FLAIR's inversion time comes from.

## 5. Spatial encoding: giving every voxel an address

### (Beginner) The orchestra

So far the coil hears one summed note from the entire head with no idea where any

of it came from. Gradients solve this.

A gradient is a small extra magnetic field that varies linearly across space.

Since precession frequency tracks field strength, a gradient makes position

determine pitch. Imagine an orchestra where the conductor forces every player to

tune according to where they sit — the left side plays flat, the right side sharp.

Now a single recording of the whole orchestra can be unmixed by pitch, and you can

work out how much sound came from each seat.

MRI uses three gradients for three jobs:

Slice select ($G_z$) — during the RF pulse, so only one slab is excited.

Frequency encode ($G_x$) — during readout, so position maps to frequency.

Phase encode ($G_y$) — briefly before readout, leaving each row with a

different phase offset.

Figure 5 — (a) The readout gradient makes frequency a linear function of position; a 10 mT/m gradient over a 240 mm field of view spreads the signal across roughly ±51 kHz. (b) The phase-encode gradient leaves a phase ramp instead. (c) Multiplying the object by one such spatial-frequency pattern and integrating gives exactly one k-space sample. (d) Slice selection picks the slab by RF centre frequency.

### (Working) Why phase encoding is the slow axis

Frequency encoding is nearly free: one readout captures a whole line of spatial

frequencies at once, because the receiver can distinguish many frequencies

simultaneously within a single acquisition window.

Phase encoding cannot do that. Phase is only measurable modulo $2\pi$ at one

instant, so each phase-encode step requires its own excitation and readout.

An image with 128 phase-encode lines therefore takes 128 repetitions, and total

scan time is roughly

$$T_{\text{scan}} \approx TR \times N_{\text{PE}} \times N_{\text{averages}}$$

This asymmetry drives nearly every acceleration technique in MRI: parallel

imaging, partial Fourier, and EPI all attack the phase-encode axis, because that

is where the time goes. It is also why distortion and blurring appear along the

phase-encode direction and essentially never along the readout direction — a fact

you will meet again in §16.

## 6. k-space: where the scanner really writes its data

### (Beginner) The recipe, not the cake

Here is the idea that makes MRI click, and it is worth slowing down for.

The scanner does not measure the image. It measures the image's ingredients.

Any picture can be built by adding together striped patterns of different

spacings, orientations and strengths — coarse stripes for the broad shapes, fine

stripes for the sharp edges. The list of "how much of each stripe pattern is

present" is a complete description of the picture: given the list, you can rebuild

the image exactly.

That list is k-space, and it is what the scanner actually fills in. A k-space

plot is not a squashed picture of a head; it is a recipe. The Fourier transform is

the oven.

Figure 6 — Top: the image and its k-space. Bottom: reconstructing from the centre alone gives a blurred image that keeps all the contrast; from the edges alone gives edges with no contrast; and skipping every second line folds the image onto itself. All three are genuine Fourier transforms of the same data.

### (Working) Reading the map

The geometry is worth memorising because it explains a great deal:

Centre of k-space = low spatial frequencies = overall brightness and

contrast. Corrupt the centre and the image is ruined.

Edges of k-space = high spatial frequencies = fine detail and sharp

boundaries. Discard them and the image blurs but stays recognisable.

Extent of k-space sampled sets spatial resolution.

Spacing between k-space samples sets the field of view.

That last pair is the one people invert, so state it plainly: how far out you go

determines how fine the detail, and how finely you step determines how wide the

picture. Step too coarsely and the image wraps around on itself, which is the

aliasing in the bottom-right panel.

### (Advanced) The signal equation

The measured signal at time $t$ is

$$S(t) = \int_{\text{object}} \rho(\mathbf{r}) \, e^{-i 2\pi \mathbf{k}(t) \cdot \mathbf{r}} \, d\mathbf{r}$$

where the k-space coordinate is the time integral of the gradient waveform:

$$\mathbf{k}(t) = \bar{\gamma} \int_0^t \mathbf{G}(\tau) \, d\tau$$

This is exactly a Fourier transform, so $\rho(\mathbf{r})$ is recovered by

inverting it. The second equation is the one to internalise for sequence design:

the gradient waveform is a steering wheel, and k-space is the terrain.

Playing a gradient moves you through k-space at a rate set by its amplitude.

Every sequence in Part III is a different route through the same terrain.

## 7. Resolution, field of view, and aliasing

### (Working) Four numbers, tightly coupled

| Quantity | Determined by | Relationship |
| --- | --- | --- |
| Field of view | k-space **sample spacing** $\Delta k$ | $\text{FOV} = 1/\Delta k$ |
| Voxel size | k-space **extent** $k_{\max}$ | $\Delta x = 1/(2k_{\max})$ |
| Matrix size | FOV ÷ voxel size | $N = \text{FOV}/\Delta x$ |
| Scan time | number of phase-encode lines | $\propto N_{\text{PE}}$ |

If anatomy extends beyond the field of view, it does not vanish — it **wraps

around** and reappears on the opposite side, because the Fourier transform of a

finitely-sampled signal is periodic. In diffusion EPI this shows up as the

subject's neck or shoulders folding into the brain, and it will happily corrupt

the parts of the image you care about.

### (Advanced) Partial Fourier

Because the object is real-valued, its k-space is (in principle) conjugate

symmetric: $S(-k) = S^*(k)$. So you could acquire just over half of k-space and

synthesise the rest. Partial Fourier does exactly this, typically sampling

6/8 or 5/8 of the lines, and it is common in diffusion imaging because it shortens

the echo train.

It is not free. Real phase errors — from motion, flow, and the very field

inhomogeneity we are about to discuss — break the symmetry assumption, so partial

Fourier reconstruction introduces blurring along the phase-encode axis and can

amplify motion artefact. It buys a shorter TE, which raises SNR; it costs

robustness. The homodyne and POCS reconstructions differ in how gracefully they

degrade when the assumption fails.

# Part III — Sequences

## 8. Spin echo and gradient echo

### (Working) The two families

Figure 7 — Spin echo (top) inserts a 180° refocusing pulse at TE/2, so the echo at TE is governed by true $T_2$. Gradient echo (bottom) refocuses with a gradient reversal instead: much faster, but field imperfections are never undone, so the echo decays with $T_2^$.*

Nearly every sequence descends from one of two ideas.

Spin echo refocuses with a 180° RF pulse. Because that pulse reverses static

dephasing, the sequence is robust to field inhomogeneity and gives true $T_2$

contrast. The cost is time (RF pulses are slow) and RF power deposition.

Gradient echo refocuses by reversing the readout gradient: dephase with a

negative lobe, then rephase with a positive one. No 180° pulse means it can run

far faster and at low flip angles, but static inhomogeneity is never undone, so

contrast follows $T_2^*$. This sensitivity is sometimes the point — susceptibility

weighted imaging and BOLD fMRI both depend on it — and sometimes the problem.

|  | Spin echo | Gradient echo |
| --- | --- | --- |
| Refocusing | 180° RF pulse | gradient reversal |
| Decay constant | $T_2$ | $T_2^*$ |
| Inhomogeneity | corrected | preserved |
| Speed | slower | faster |
| SAR | higher | lower |
| Typical use | $T_2$w, FLAIR, DWI readout | MPRAGE, SWI, BOLD |

### (Advanced) Turbo spin echo

A single 180° pulse can be replaced by a train of them, collecting a different

phase-encode line after each. This is turbo/fast spin echo, and the echo train

length is the acceleration factor. The catch is that successive echoes have

experienced different amounts of $T_2$ decay, so different parts of k-space carry

different weightings. Since contrast is dominated by the k-space centre, the

effective TE is whichever echo happens to land there. This is why two TSE

protocols with identical nominal TE can look different: what matters is the view

ordering, not the label.

## 9. Echo-planar imaging: the whole plane in one shot

### (Beginner) One breath, the whole song

Conventional imaging collects one k-space line per excitation. Diffusion imaging

cannot afford that: you need dozens of volumes, and the subject would have to hold

still for an hour.

EPI takes the extreme option. After a single excitation, it oscillates the

readout gradient back and forth as fast as the hardware allows, zig-zagging

through all of k-space before the signal dies. A whole slice in roughly 50 ms;

a whole brain in a couple of seconds.

The analogy: instead of singing one note per breath, EPI sings the entire song on

a single breath. It is fast, and it works — but the singer is running out of air

throughout, and you can hear it in the last few bars.

### (Working) What running out of air costs

Figure 8 — (a) The zig-zag trajectory: readout sweeps back and forth while small "blips" step the phase-encode position. (b) $T_2^$ decay across a 38 ms echo train. (c) Because later k-space lines are systematically weaker, the effective k-space filter blurs the image along the phase-encode axis.*

Three consequences follow directly, and all three are visible in the data this

pipeline processes:

Blurring along the phase-encode axis. Signal decays during the train, so

high-$k_y$ lines are attenuated. That is a low-pass filter in one direction

only.

Geometric distortion. The phase-encode axis accumulates phase error for the

entire readout duration, so off-resonance turns into large spatial

displacement. This is §16, and it is the big one.

Ghosting. Alternate lines are collected under opposite gradient polarity.

Any timing mismatch between them produces a copy of the image shifted by half

the field of view — the N/2 ghost.

The single most useful protocol lever here is shortening the echo train, via

parallel imaging, partial Fourier, or multi-shot EPI. Everything improves at once.

## 10. Accelerating: parallel imaging and multiband

### (Working) Two different kinds of cheating

Parallel imaging (SENSE, GRAPPA) skips phase-encode lines and recovers the

missing information from the fact that multiple receive coils each see the head

with a different spatial sensitivity. Undersampling alone would fold the image

(Figure 6, bottom right), but the coil sensitivity profiles provide the extra

constraint needed to unfold it.

Acceleration factor $R$: scan time and echo train shrink by $R$.

SNR cost: $\text{SNR} \propto 1/(g\sqrt{R})$, where the g-factor $g \geq 1$

depends on coil geometry and worsens toward the centre of the head.

Typical diffusion use: $R = 2$. Higher factors produce central noise

amplification and residual unfolding artefacts.

Multiband (simultaneous multi-slice) is a different trick: excite several

slices at once with a composite RF pulse and separate them afterwards using coil

sensitivity plus deliberately applied inter-slice phase shifts (blipped CAIPI).

Multiband factor MB: slices per unit time improve by MB.

The crucial difference from parallel imaging: it does not shorten the echo

train, so it does not reduce distortion — and, because it does not

undersample within a slice, its intrinsic SNR penalty is much smaller.

Cost: slice leakage between simultaneously excited slices, and higher SAR.

The two combine, and modern diffusion protocols routinely use both. If you are

reading a protocol and want to predict distortion, look at the in-plane

acceleration and partial Fourier, not the multiband factor.

## 11. The structural scan: MPRAGE and FLAIR

### (Working) Why the pipeline needs a good $T_1$

This pipeline's anatomical steps — FreeSurfer or FastSurfer, the 5TT image, the

parcellation — all rest on one $T_1$-weighted volume. Its quality sets a ceiling

on everything downstream, and no amount of clever processing recovers a

poorly-acquired structural scan.

MPRAGE is a 3-D gradient echo preceded by an inversion pulse. The inversion is

what generates strong grey/white contrast, and the 3-D readout gives isotropic

sub-millimetre voxels with good SNR. Typical 3 T parameters: TI ≈ 900 ms,

TR ≈ 2300 ms, flip angle 8–9°, 1 mm isotropic.

What matters for surface reconstruction:

| Property | Why it matters |
| --- | --- |
| Isotropic ~1 mm | the cortical ribbon is only 2–4 mm thick; anisotropic voxels blur it |
| Good GM/WM contrast | the white surface is placed on the intensity gradient |
| Minimal motion | motion blurs precisely the boundary being detected |
| Bias field within reason | correctable by N4, but not without limit |

FLAIR adds a long inversion (~2.5 s) to null CSF, leaving lesions conspicuous.

It is not required by this pipeline, but it is the sequence that makes white

matter hyperintensities visible, and those matter for interpreting connectivity

findings in older or vascular cohorts (see brain.md §12).

# Part IV — Diffusion

## 12. Brownian motion and what diffusion MRI measures

### (Beginner) Ink in water, and ink in a bundle of straws

Drop ink into still water and it spreads on its own. Nothing pushes it; water

molecules are in constant thermal motion, and random jostling carries the ink

outward. That is Brownian motion, and every water molecule in the brain is

doing it, all the time.

The key insight of diffusion MRI is that this random motion is **not random in its

geometry**. Where water is free, it spreads equally in all directions. Where it is

hemmed in, it cannot.

Now imagine dropping ink into a tightly packed bundle of drinking straws. Along

the straws, it spreads freely. Across them, it is blocked by the walls. Measure

how far the ink got in various directions and you can deduce which way the straws

point — without ever seeing a straw.

White matter is that bundle of straws. Axons, wrapped in myelin and packed in

parallel, restrict water movement across the fibre far more than along it.

Diffusion MRI measures displacement in many directions and infers fibre

orientation from the pattern. Tractography then joins those local orientations

into pathways, and this pipeline counts the pathways between regions to build a

connectome.

Worth being clear about what is and is not measured: we never see an axon.

Diffusion MRI voxels are 1–3 mm across and axons are microns. We observe the

average behaviour of water in a volume containing perhaps a million axons and

infer a dominant orientation. That inferential gap is the origin of every

tractography controversy in the literature.

### (Working) Numbers

Free water at body temperature has a diffusion coefficient of about

$3.0 \times 10^{-3}$ mm²/s. In tissue, apparent values are much lower:

| Tissue | Apparent diffusivity (mm²/s) | Anisotropy |
| --- | --- | --- |
| CSF | ~3.0 × 10⁻³ | none |
| Grey matter | ~0.8 × 10⁻³ | low |
| White matter (along fibre) | ~1.7 × 10⁻³ | high |
| White matter (across fibre) | ~0.3 × 10⁻³ | — |

In roughly 50 ms of diffusion time, water diffuses on the order of 10 µm — a

distance comparable to cell dimensions, which is exactly why the measurement is

sensitive to microstructure it cannot resolve.

Acute stroke is the clinical showcase: cytotoxic oedema traps water inside

swollen cells, diffusivity plummets, and the lesion lights up on DWI within

minutes, long before anything appears on CT or $T_2$.

## 13. The Stejskal–Tanner experiment and the b-value

### (Beginner) Label, wait, re-read

The experiment, invented by Stejskal and Tanner in 1965, is beautifully simple:

Label every spin with a position-dependent phase, using a gradient pulse.

Think of stamping each water molecule with its map reference.

Wait a short interval $\Delta$, during which molecules diffuse.

Re-read with an equal and opposite gradient pulse. Anything that stayed

put gets its phase perfectly cancelled. Anything that moved does not.

Spins that stayed put add together coherently and give full signal. Spins that

moved have scattered phases and partially cancel. **Signal loss is the

measurement.** More movement means less signal.

Figure 9 — (a) The PGSE sequence: two matched gradient pulses of duration $\delta$ straddling the 180° refocusing pulse, separated by the diffusion time $\Delta$. (b) Static spins are perfectly rephased; diffusing spins are not, and scattered phases cancel. (c) Attenuation against b-value. CSF, diffusing ten times faster than white matter, is almost entirely suppressed by b = 1000 — which is why ventricles are black on a diffusion-weighted image.

### (Working) The b-value

How strongly you diffusion-weight is summarised in one number:

$$b = \gamma^2 G^2 \delta^2 \left(\Delta - \frac{\delta}{3}\right)$$

with $G$ the gradient amplitude, $\delta$ its duration, $\Delta$ the separation.

Units are s/mm². Signal attenuates as

$$S = S_0 \, e^{-bD}$$

| b-value | Effect | Use |
| --- | --- | --- |
| 0 | none | reference volume, needed to normalise |
| 700–1000 | moderate | clinical DWI, DTI, this pipeline |
| 2000–3000 | strong | crossing fibres, HARDI, CSD |
| >3000 | severe | microstructure models; low SNR |

The trade-off is unavoidable: raising b increases angular contrast but reduces

signal exponentially, so noise eventually dominates. And note what $b$ hides —

very different combinations of $G$, $\delta$ and $\Delta$ give the same $b$ but

probe different length scales. Two protocols quoting "b = 1000" are not

necessarily measuring the same thing, which is a real obstacle to multi-site

harmonisation.

### (Advanced) Why the $-\delta/3$

The idealised derivation assumes infinitesimally short gradient pulses, so that

all diffusion happens in the interval $\Delta$ between them. Real pulses have

finite duration $\delta$, during which spins are both being labelled and

diffusing. Integrating the phase properly over the finite pulses yields the

correction, giving an effective diffusion time of $\Delta - \delta/3$. For

typical clinical parameters ($\delta \approx 30$ ms, $\Delta \approx 50$ ms) this

is not a small correction — it changes the effective diffusion time by about 20%.

## 14. Directions, shells, and the gradient table

### (Working) How many directions, and where

*Figure 10 — (a) Signal as a function of gradient direction. Isotropic tissue gives a circle; anisotropic tissue gives a peanut, with least signal along the fibre, because that is the direction water travels furthest. (b) A well-designed 60-direction table spreads samples evenly over the sphere, exploiting the fact that $+\mathbf{g}$ and $-\mathbf{g}$ are equivalent. (c) The b-value list volume by volume, with b = 0 images interleaved throughout for motion tracking.*

Diffusion is measured one direction at a time, so the gradient table is a design

choice with real consequences:

| Directions | Supports |
| --- | --- |
| 6 | the diffusion tensor, minimally — no redundancy, no error estimate |
| 30 | robust DTI: FA, MD, and a stable principal direction |
| 60+ | constrained spherical deconvolution, crossing fibres |
| 90+ | higher-order models, multi-shell microstructure |

Directions should be spread as evenly as possible over the sphere; the standard

construction minimises an electrostatic repulsion energy between antipodal point

pairs. Two details that bite in practice:

Antipodal symmetry. Diffusion cannot distinguish $+\mathbf{g}$ from

$-\mathbf{g}$, so directions live on a half-sphere. A table that fails to

account for this wastes measurements.

b = 0 volumes must be interleaved, not all collected at the start. They are

the anchors for motion and eddy-current correction, and anchors are only useful

if they are distributed through the series.

Single-shell versus multi-shell is the choice that determines which

reconstruction models are available. A single shell (this pipeline: b = 1000 plus

b = 0) supports DTI and single-shell CSD variants. Multi-shell adds b = 2000 and

b = 3000, enabling multi-tissue CSD and microstructure models such as NODDI, at

the cost of proportionally more scan time.

This is precisely why the pipeline's QSIRecon spec uses SS3T — single-shell

three-tissue CSD. Standard multi-tissue CSD needs multiple shells to separate the

tissue response functions; SS3T was developed to achieve a comparable separation

from a single shell. The reconstruction choice is downstream of an acquisition

choice made years earlier, and cannot be revisited afterwards. See

pipeline_science.md for what SS3T then does with it.

## 15. Designing a diffusion protocol

### (Working) A worked example

Suppose you have 10 minutes for diffusion in a 45-minute session, and the goal is

structural connectomics.

| Decision | Choice | Reasoning |
| --- | --- | --- |
| b-value | 1000 s/mm² | good SNR; adequate for CSD-family models |
| Directions | 60 | supports spherical deconvolution; ~30 would cap you at DTI |
| b = 0 volumes | 6–7, interleaved | anchors for motion and eddy correction |
| Voxel size | 2 mm isotropic | 1.5 mm would cost over half the SNR |
| Partial Fourier | 6/8 | shortens TE, hence more signal |
| In-plane accel | R = 2 | shortens echo train, hence less distortion |
| Multiband | 2–3 | more slices per second, little SNR penalty |
| **Reversed PE** | **a few b = 0 volumes, opposite direction** | **enables proper distortion correction** |

The last row is the one most often omitted and the most expensive to omit. A

handful of reversed-phase-encode b = 0 volumes costs under a minute and is the

difference between principled distortion correction and a fieldmap-free

approximation. If you can change one thing about a legacy protocol, change this.

Total volumes: 66. At TR ≈ 4 s with multiband, roughly 4–5 minutes, leaving room

for the reversed-PE acquisition and a repeat if the first attempt fails QC.

# Part V — Everything that goes wrong

## 16. Susceptibility distortion and the reversed-PE trick

### (Beginner) Where the head stops being a sphere of water

Magnetic susceptibility describes how a material distorts a magnetic field.

Air and tissue differ substantially, so wherever they meet — the sinuses behind

the nose, the air cells in the temporal bone, the ear canals — the field is

locally warped no matter how well the scanner is shimmed.

Now recall the chain of logic from §1 and §5. Field determines frequency.

Frequency determines apparent position. So a field error becomes a position

error: signal is reconstructed in the wrong place. Orbitofrontal and anterior

temporal cortex, which sit directly against air cavities, are exactly where this

is worst — and both are of central interest in traumatic brain injury.

### (Working) Why the phase-encode axis, and how far

Figure 11 — Off-resonance near air/bone interfaces (top left) displaces signal along the phase-encode axis. The AP and PA acquisitions distort in exactly opposite senses — the red outline marks true anatomy, and the frontal region piles up in one and stretches in the other. Bottom left: displacement in voxels. Bottom middle: the geometry recovered from the pair. Bottom right: halving the readout time halves the distortion.

The displacement is given by a strikingly simple relation:

$$\Delta y \; [\text{voxels}] = \Delta f \; [\text{Hz}] \times T_{\text{readout}} \; [\text{s}]$$

That is the whole story, and two things follow.

First, distortion is confined to the phase-encode axis. The readout axis

samples an entire line in under a millisecond, so accumulated phase error is

negligible. The phase-encode axis accumulates error across the whole echo train,

tens of milliseconds. The ratio is roughly the echo train length, so distortion is

about a hundred times worse in one direction than the other.

Second, it scales linearly with readout duration. A 200 Hz offset with a 50 ms

readout displaces signal by 10 voxels — 2 cm at 2 mm resolution. Halve the readout

via parallel imaging and you halve the displacement.

The reversed-PE trick. Reverse the phase-encode direction and every

displacement reverses sign while keeping its magnitude. One image compresses

exactly where the other stretches. Given both, the underlying field map (and hence

the true geometry) can be estimated. This is what FSL's topup does, and what

QSIPrep's SDC stage invokes. It requires that the reversed-PE data was acquired —

which is why §15 insisted on it, and why the sidecar fields in §19 must be right.

### (Advanced) When you do not have a reversed pair

Fallbacks exist, in descending order of trustworthiness:

Acquired fieldmap (phase-difference or multi-echo GRE): measures

off-resonance directly. Good, though it needs unwrapping and is itself

sensitive to motion between the fieldmap and the diffusion scan.

Fieldmap-less SDC (--use-syn-sdc): non-linearly registers the distorted

EPI to an undistorted $T_1$, constraining the warp to the phase-encode axis.

Surprisingly effective, but it is inference from anatomy, not measurement,

and it degrades exactly where anatomy is abnormal — which in a lesion cohort

is the region of interest.

Nothing: leaves centimetre-scale misplacement in frontal and temporal

cortex, which then silently misassigns streamlines to the wrong parcels.

This pipeline's behaviour is documented in fmaps.md, and the

sidecar requirements in bids.md. The relevant point for

acquisition is that option 1 and option 2 are decided at the scanner, not later.

## 17. Eddy currents, motion, and the rest of the artefact zoo

### (Working) The gallery

Figure 12 — Each panel starts from the same image and corrupts the k-space data in the way the scanner does: truncating for Gibbs, randomising line phase for motion, shearing for eddy currents, adding a single bad sample for the spike, mismatching odd and even lines for the N/2 ghost, and adding Rician noise.

| Artefact | Mechanism | Appearance | Handled by |
| --- | --- | --- | --- |
| **Eddy currents** | switching diffusion gradients induces currents in conductors, which add unwanted fields | shear, scale and shift that vary *with gradient direction* | `eddy` |
| **Motion** | subject moves between or during volumes | ghosting, blurring, signal dropout | `eddy`, volume-to-volume registration |
| **Gibbs ringing** | k-space truncated at finite extent | ripples adjacent to sharp edges | subvoxel-shift unringing |
| **Spike / RF noise** | a single corrupt k-space sample | stripes across the whole image | QC and volume rejection |
| **N/2 ghost** | odd/even echo timing mismatch in EPI | copy shifted half a field of view | reconstruction-side phase correction |
| **Thermal noise** | receiver electronics, patient | grain; Rician, not Gaussian | denoising (MP-PCA) |

Two of these deserve emphasis because they interact with diffusion specifically.

Eddy currents are direction-dependent. Each diffusion direction uses a

different gradient combination, so each volume is distorted differently. Rigid

registration cannot fix it; you need a model that accounts for the gradient

direction, which is what FSL's eddy provides. Uncorrected, it appears as

apparent anisotropy that is actually an artefact of misregistration.

Noise in magnitude images is Rician, not Gaussian. Taking the magnitude of

complex data with Gaussian noise rectifies it, so the noise floor sits above

zero. At high b-values, where true signal approaches the floor, this biases

diffusivity estimates downward. It is a systematic bias, not random error, and it

is why denoising before any model fitting is the recommended order of operations.

## 18. Noise, SNR, and the trade-off triangle

### (Beginner) You cannot have all three

*Figure 13 — (a) SNR scales with voxel volume, so shrinking the edge from 2.5 mm to 1.5 mm costs nearly five-fold. (b) Averaging helps only as $\sqrt{N}$, so four times the scan time buys twice the SNR. (c) Resolution, SNR and speed: pick two.*

Every protocol argument reduces to three quantities that cannot be maximised

together:

$$\text{SNR} \propto (\Delta x \, \Delta y \, \Delta z) \times \sqrt{N_{\text{avg}} \, T_{\text{acq}}}$$

Resolution is brutally expensive: SNR follows voxel volume, so halving the

edge length costs a factor of eight. Averaging is inefficient: signal adds

coherently and noise adds in quadrature, so SNR improves only as the square root

of the time spent. Four times the scan for double the SNR is a poor exchange when

subjects are moving.

The practical consequence is that resolution should be set by the smallest

structure you genuinely need to resolve, not by ambition. For structural

connectomics, 2 mm isotropic is the common compromise: fine enough to follow major

tracts, coarse enough to keep SNR workable at b = 1000 within a clinical session.

# Part VI — From scanner to pipeline

## 19. DICOM to BIDS: what must survive the conversion

### (Working) Metadata is not optional

Figure 14 — The full chain. Acquisition decisions on the left propagate all the way to the connectome on the right, and several of them cannot be undone afterwards.

The scanner exports DICOM; the pipeline consumes BIDS.

dcm2niix bridges them, producing a NIfTI image, a .bval/.bvec pair, and a

JSON sidecar. Everything the pipeline knows about how the data was acquired comes

from that sidecar, and wrong metadata is worse than missing metadata: missing

values usually stop the pipeline, whereas wrong ones let it run confidently to a

wrong answer.

| Field | Meaning | Consequence if wrong |
| --- | --- | --- |
| `PhaseEncodingDirection` | axis and polarity (`j`, `j-`, …) | distortion correction applied **backwards**, doubling the error |
| `TotalReadoutTime` | echo train duration | distortion scaled incorrectly |
| `EffectiveEchoSpacing` | time between echoes | same |
| `EchoTime`, `RepetitionTime` | timings | modelling and QC |
| `IntendedFor` | which scans a fieldmap applies to | fieldmap silently ignored |
| `SliceTiming` | acquisition order | relevant for slice-wise outlier detection |
| `.bvec` | gradient directions | **wrong fibre orientations, hence wrong tractography** |
| `.bval` | b-values | wrong model fitting |

The .bvec file deserves particular caution. Gradient directions are expressed in

a coordinate frame, and the conventions differ between vendors, between

dcm2niix versions, and between image orientations. A left–right flip produces a

tractogram that looks entirely plausible and is systematically wrong. The standard

sanity check is to reconstruct a tensor and confirm that the corpus callosum comes

out left–right (red in the usual colour convention) and the corticospinal tract

superior–inferior (blue). Do this once per protocol, not once per study.

Repairing sidecars is a pre-pipeline step here, deliberately kept outside the

processing so that the fix is explicit, inspectable and version-controlled. See

bids.md.

## 20. How acquisition choices reach the connectome

### (Working) The causal chain

Worth stating plainly, because it is easy to treat acquisition as somebody else's

problem:

| Acquisition choice | Immediate effect | Consequence for the connectome |
| --- | --- | --- |
| Single shell at b = 1000 | multi-tissue CSD unavailable | forces SS3T; multi-shell models permanently out of reach |
| 60 directions | supports spherical deconvolution | fewer would restrict you to DTI and lose crossing fibres |
| 2 mm isotropic | moderate SNR | smaller voxels would raise noise and destabilise tracking |
| No reversed-PE volumes | no `topup` | frontal/temporal misplacement, streamlines assigned to wrong parcels |
| Long echo train | large distortion | same, in proportion to readout time |
| Motion during the scan | volume-wise corruption | dropped volumes, or a subject failing QC entirely |
| $T_1$ resolution and contrast | surface placement quality | parcel boundaries, hence every matrix entry |

The recurring theme is that acquisition errors do not announce themselves. A

distorted brain still produces a full connectivity matrix of exactly the right

size. Nothing in the pipeline will refuse to run. The matrix will simply be wrong

in a spatially structured way — worst in orbitofrontal and temporal regions,

which is precisely where a TBI study is looking.

## 21. Acquisition QC: what to check before you process

### (Working) A checklist

Cheap to run, and each one catches a failure that is expensive later.

On the structural scan

[ ] Grey/white contrast adequate, and inversion preparation actually used

[ ] No motion blur at the grey/white boundary

[ ] Full brain coverage including cerebellum and vertex

[ ] Bias field within N4's ability to correct

On the diffusion series

[ ] Volume count matches the expected number

[ ] .bval/.bvec row counts match the number of volumes

[ ] b = 0 volumes present and interleaved, not front-loaded

[ ] Reversed-PE volumes present, if the protocol specified them

[ ] No wrap-around of neck or shoulders into the brain

[ ] Signal dropout absent in temporal lobes and cerebellum

[ ] Tensor colour map shows callosum left–right, corticospinal superior–inferior

On the metadata

[ ] PhaseEncodingDirection present, and its polarity verified against the pair

[ ] TotalReadoutTime present and physically plausible (tens of ms)

[ ] IntendedFor resolves to files that actually exist

[ ] Parameters consistent across subjects in the cohort

The last box is the one that quietly ruins group analyses. A protocol changed

midway through recruitment — a different readout time, a different multiband

factor — introduces a systematic difference that no amount of downstream

harmonisation fully removes. Record the protocol, check it periodically, and

treat any change as a covariate at minimum.

## Glossary

ADC — apparent diffusion coefficient; measured diffusivity, "apparent" because

tissue restriction makes it lower than free water.

b-value — diffusion weighting strength, s/mm².

BIDS — Brain Imaging Data Structure; the directory and metadata convention.

CSD — constrained spherical deconvolution; recovers fibre orientation

distributions from diffusion signal.

EPI — echo-planar imaging; fills k-space after a single excitation.

Ernst angle — flip angle maximising signal for a given TR and $T_1$.

FLAIR — fluid-attenuated inversion recovery; nulls CSF.

Flip angle — how far an RF pulse tips the magnetisation.

FOV — field of view; the spatial extent imaged.

GRAPPA / SENSE — parallel imaging methods using multiple coil sensitivities.

Gibbs ringing — ripples near sharp edges from truncating k-space.

g-factor — geometry-dependent noise amplification in parallel imaging.

k-space — the spatial-frequency domain the scanner actually samples.

Larmor frequency — precession frequency, $\bar{\gamma}B_0$.

MPRAGE — inversion-prepared 3-D gradient echo; the standard structural $T_1$w.

Multiband — exciting several slices simultaneously.

N/2 ghost — half-FOV image copy from odd/even echo mismatch in EPI.

Partial Fourier — acquiring just over half of k-space and exploiting symmetry.

PGSE — pulsed-gradient spin echo; the Stejskal–Tanner diffusion sequence.

Phase encoding — spatial encoding by phase; the slow, distortion-prone axis.

Rician noise — the noise distribution of magnitude MRI data; biased above zero.

SAR — specific absorption rate; RF power deposition limit.

SDC — susceptibility distortion correction.

Shim — correction coils that flatten $B_0$ inhomogeneity.

SNR — signal-to-noise ratio.

SS3T — single-shell three-tissue CSD.

Susceptibility — a material's response to a magnetic field; mismatches at

air/tissue interfaces cause distortion.

TE — echo time; controls $T_2$ weighting.

TI — inversion time; controls which tissue is nulled.

topup — FSL tool estimating the field from reversed-PE image pairs.

TR — repetition time; controls $T_1$ weighting.

$T_1$ — longitudinal relaxation time.

$T_2$ — transverse relaxation time.

**$T_2^*$** — transverse relaxation including static inhomogeneity; shorter.

## References

### MRI physics: textbooks and general reference

Nishimura DG. Principles of Magnetic Resonance Imaging. Stanford University, 2010. — the clearest treatment of k-space and signal encoding.

Bernstein MA, King KF, Zhou XJ. Handbook of MRI Pulse Sequences. Elsevier, 2004. — the standard reference for sequence design.

Haacke EM, Brown RW, Thompson MR, Venkatesan R. Magnetic Resonance Imaging: Physical Principles and Sequence Design. 2nd ed. Wiley, 2014.

Brown RW, Cheng YCN, Haacke EM, et al. MRI: Physical Principles and Sequence Design. Wiley-Blackwell.

Elster AD. Questions and Answers in MRI — mriquestions.com. Free, searchable, and unusually good at explaining why.

McRobbie DW, Moore EA, Graves MJ, Prince MR. MRI from Picture to Proton. 3rd ed. Cambridge University Press, 2017. — the friendliest starting point.

### Foundational papers

Bloch F. Nuclear induction. Physical Review 1946;70:460–474.

Hahn EL. Spin echoes. Physical Review 1950;80(4):580–594.

Lauterbur PC. Image formation by induced local interactions: examples employing nuclear magnetic resonance. Nature 1973;242:190–191.

Mansfield P. Multi-planar image formation using NMR spin echoes. Journal of Physics C 1977;10(3):L55–L58. (EPI)

Ljunggren S. A simple graphical representation of Fourier-based imaging methods. Journal of Magnetic Resonance 1983;54(2):338–343. (k-space formalism)

### Diffusion MRI

Stejskal EO, Tanner JE. Spin diffusion measurements: spin echoes in the presence of a time-dependent field gradient. Journal of Chemical Physics 1965;42:288–292.

Le Bihan D, Breton E, Lallemand D, et al. MR imaging of intravoxel incoherent motions. Radiology 1986;161(2):401–407.

Basser PJ, Mattiello J, LeBihan D. MR diffusion tensor spectroscopy and imaging. Biophysical Journal 1994;66(1):259–267.

Jones DK (ed). Diffusion MRI: Theory, Methods, and Applications. Oxford University Press, 2010.

Jones DK, Knösche TR, Turner R. White matter integrity, fiber count, and other fallacies: the do's and don'ts of diffusion MRI. NeuroImage 2013;73:239–254. — essential reading before interpreting any diffusion metric.

Tournier JD, Calamante F, Connelly A. Robust determination of the fibre orientation distribution in diffusion MRI. NeuroImage 2007;35(4):1459–1472. (CSD)

Dhollander T, Raffelt D, Connelly A. Unsupervised 3-tissue response function estimation from single-shell or multi-shell diffusion MR data. ISMRM Workshop, 2016. (SS3T)

Jones DK, Horsfield MA, Simmons A. Optimal strategies for measuring diffusion in anisotropic systems by magnetic resonance imaging. Magnetic Resonance in Medicine 1999;42(3):515–525. — where direction schemes come from.

### Artefacts and correction

Jezzard P, Balaban RS. Correction for geometric distortion in echo planar images from $B_0$ field variations. Magnetic Resonance in Medicine 1995;34(1):65–73.

Andersson JLR, Skare S, Ashburner J. How to correct susceptibility distortions in spin-echo echo-planar images: application to diffusion tensor imaging. NeuroImage 2003;20(2):870–888. (topup)

Andersson JLR, Sotiropoulos SN. An integrated approach to correction for off-resonance effects and subject movement in diffusion MR imaging. NeuroImage 2016;125:1063–1078. (eddy)

Veraart J, Novikov DS, Christiaens D, et al. Denoising of diffusion MRI using random matrix theory. NeuroImage 2016;142:394–406. (MP-PCA)

Kellner E, Dhital B, Kiselev VG, Reisert M. Gibbs-ringing artifact removal based on local subvoxel-shifts. Magnetic Resonance in Medicine 2016;76(5):1574–1581.

Gudbjartsson H, Patz S. The Rician distribution of noisy MRI data. Magnetic Resonance in Medicine 1995;34(6):910–914.

Tustison NJ, Avants BB, Cook PA, et al. N4ITK: improved N3 bias correction. IEEE Transactions on Medical Imaging 2010;29(6):1310–1320.

### Acceleration

Pruessmann KP, Weiger M, Scheidegger MB, Boesiger P. SENSE: sensitivity encoding for fast MRI. Magnetic Resonance in Medicine 1999;42(5):952–962.

Griswold MA, Jakob PM, Heidemann RM, et al. Generalized autocalibrating partially parallel acquisitions (GRAPPA). Magnetic Resonance in Medicine 2002;47(6):1202–1210.

Setsompop K, Gagoski BA, Polimeni JR, et al. Blipped-controlled aliasing in parallel imaging for simultaneous multislice EPI with reduced g-factor penalty. Magnetic Resonance in Medicine 2012;67(5):1210–1224. (blipped-CAIPI multiband)

### Protocol design, standards, and reproducibility

Gorgolewski KJ, Auer T, Calhoun VD, et al. The brain imaging data structure. Scientific Data 2016;3:160044. (BIDS)

Li X, Morgan PS, Ashburner J, et al. The first step for neuroimaging data analysis: DICOM to NIfTI conversion. Journal of Neuroscience Methods 2016;264:47–56. (dcm2niix)

Cieslak M, Cook PA, He X, et al. QSIPrep: an integrative platform for preprocessing and reconstructing diffusion MRI data. Nature Methods 2021;18:775–778.

Glasser MF, Sotiropoulos SN, Wilson JA, et al. The minimal preprocessing pipelines for the Human Connectome Project. NeuroImage 2013;80:105–124.

Nichols TE, Das S, Eickhoff SB, et al. Best practices in data analysis and sharing in neuroimaging using MRI. Nature Neuroscience 2017;20:299–303.

Fortin JP, Parker D, Tunç B, et al. Harmonization of multi-site diffusion tensor imaging data. NeuroImage 2017;161:149–170.

### Tissue property values used in the figures

Wansapura JP, Holland SK, Dunn RS, Ball WS. NMR relaxation times in the human brain at 3.0 Tesla. Journal of Magnetic Resonance Imaging 1999;9(4):531–538.

Stanisz GJ, Odrobina EE, Pun J, et al. T1, T2 relaxation and magnetization transfer in tissue at 3T. Magnetic Resonance in Medicine 2005;54(3):507–512.

Bojorquez JZ, Bricq S, Acquitter C, et al. What are normal relaxation times of tissues at 3 T? Magnetic Resonance Imaging 2017;35:69–80.

## Reproducing the figures

Every figure is computed from the equations, with no input data of any kind:

python3 dwi_pipeline/scripts/make_acquisition_figures.py \
    --out-dir dwi_pipeline/figures/acquisition

Requires numpy, scipy and matplotlib. Runs in a few seconds.

Provenance

| Figure | How it is produced | Measured data? |
| --- | --- | --- |
| 1 | $f_0 = \bar\gamma B_0$ evaluated over field strength | No — exact physics |
| 2 | Rotating-frame trajectory; $\sin/\cos\alpha$; slice profile from bandwidth ÷ gradient | No |
| 3 | Bloch solutions at published 3 T constants (refs 36–38) | No — literature constants |
| 4 | Signal equations of §4 applied to an analytic phantom | No |
| 5 | Gradient-induced frequency and phase maps, 10 mT/m over 240 mm | No |
| 6 | Genuine 2-D FFT of the phantom, variously filtered | No |
| 7 | Schematic timing; echo envelopes are exponentials | No |
| 8 | $T_2^*$ decay applied as a k-space filter along $k_y$ | No |
| 9 | Stejskal–Tanner relation with literature diffusivities | No |
| 10 | $S = S_0 e^{-b\mathbf{g}^T\mathbf{D}\mathbf{g}}$; Fibonacci-sphere directions | No |
| 11 | Displacement $\Delta y = \Delta f \times T_{\text{readout}}$ applied as a warp | No |
| 12 | Each artefact induced by its real mechanism in k-space | No |
| 13 | SNR $\propto$ voxel volume $\times \sqrt{N}$ | No |
| 14 | Schematic | No |

The phantom is defined analytically in brain_phantom(): nested ellipses for the

brain, ventricles and deep nuclei, an undulating cortical ribbon of roughly

constant thickness, and explicitly cut sulcal clefts lined by cortex. It is not

derived from any scan, and nothing in this document depicts a person.

This document supports research engineering and is not a clinical reference. Tissue property values are representative literature figures at 3 T and vary with sequence, temperature, and measurement method.